| Equipment: | Wireless Blood Refrigerator Monitoring. |
| --- | --- |
| Department: | Blood Bank |
| Date: | November, 2007 |
| Written by: |  |

| #### Review and approval |  |  |  |  |
| --- | --- | --- | --- | --- |
| Name | Designation | Signature | Date | Function |

| Document History |  |  |  |
| --- | --- | --- | --- |
| Version | Issue Date | Section | Reason for Issue/Change |
| 1 | 01/11/2007 | N/A | New Document |

Contents

1	Scope

  - Objectives

  - Benefits

- Terminology

- Regulations

- Use

  - Capacity

  - Operational requirements

  - Process requirements

  - Process control

- Function

  - Operation

  - Control system requirements

  - Power Failure/Recovery

  - Recovery methods

  - Alarms/Warnings

  - User interfaces

  - Environment

  - Physical conditions

  - Environmental classifications

  - Biohazard level

  - Cleaning requirements

  - Sterilisation requirements

  - Intended operating environment

- Requirements for Installation

  - Calibration

  - Deliverables

  - Documentation

  - Training

  - Tools and Spares

  - PPM/Calibration/Servicing/Re-Validation(Mapping)

1	Scope

This document has been created to define the user requirements for the purchase of a wireless blood fridge monitoring system that will allow us to ensure that all blood component storage areas (fridges, freezers etc) across XXXX and YYYY have ‘core’ temperature alarms set between 2-6C that are monitored 24/7 and would be acted upon immediately.

1.1	Objectives

Blood components will be stored in appropriate storage conditions monitored and alarmed to specifications laid down by law and enforced by the MHRA.

  - Benefits

The purchase of this equipment is in order to correct a major non compliance following our recent MHRA inspection.

2	Terminology

This document refers to the features and attributes the monitoring system should have and how it should perform from the user perspective. The URS describes what the proposed system will do, but nothing about how it will be built or implemented.

Narrative descriptions and diagrams can be used to show the services the proposed system will deliver and its limitations.

The document can be used as a “contract” between the client and the equipment supplier who will provide and install the system.

Key words used within the URS include “must”, “must not”, “should”, “should not”, “recommended”, and “may” or “optional” and are to be interpreted as follows:

Must	this word means that the definition is an absolute requirement of the specification.

Must not 	This phrase means that the definition is an absolute prohibition of the specification.

Should	This word or the adjective “Recommended” means that there may exist valid reasons in particular circumstances to ignore a particular item, but the full implications must be understood and carefully weighed before choosing a different course.

Should not	This phrase means that there may exist valid reasons in particular circumstances when the particular behaviour is acceptable or even useful but the full implications should be understood and the case carefully weighed.

May	This word, or the adjective “optional” means that an item is truly optional

Definition of Signatures

Reviewed by

A review signature verifies that, as a competent reviewer, it is considered that the content of this document is sufficient to meet SNBTS and regulatory guidelines on validation and GXP requirements.

Approved by:

An approval signature indicates that the content of this document is sufficient to meet our and regulatory guidelines on validation and GXP. It will also ensure that the appropriate personnel have reviewed this document. The Transfusion Laboratory Manager will approve this document.

3	Regulations

- The Blood Safety and Quality Regulation (No 50) 2005 as amended.

- BS 4376: Part 1 (1991) and Part 2 (1996) : Specification for electrically operated blood storage refrigerators and freezers

- European Blood Directive-  (2002/98/EC)

### 4	Use

Blood components must be kept in an environment which will maintain them within the temperature range as defined within The Blood Safety and Quality Regulation (No 50) 2005 as amended. The blood component storage units will have its temperature monitored constantly by this new temperature monitoring system.

These component storage units hold various components either as stock in the laboratory or as allocated components issued for individual patients awaiting collection and transportation to the patient’s location.

### 4.1	Capacity

The system must be able to monitor all blood component storage devices around the Trust. This currently comprises 12 units at XXXX and 5 units at YYYY.

The system should be configurable to allow expansion of the system by addition of future temperature controlled units (indicative cost per extra unit would be useful).

4.2	Operational Requirements

- Core temperatures in each unit must be monitored at temperatures laid down by legislation and noted in section 4.3.

- Monitoring must be automatic, continuous in real time with logging at an interval of no more than 15 minutes.

- The location of each probe must be uniquely identifiable

- Alarms or alerts must be recorded in a format/medium which can be stored and retrievable for a minimum of 15 years. Also need to ensure this cannot be overwritten accidentally

- Alarms or alerts must be recorded with ability to add user notation for each incident.

- System must be able to generate automatic/customised reports and action reports/event logs.

4.3	Process Requirements

Red cell blood components must be stored with a core temperature of between +2 oC and +6oC throughout its shelf-life.

Frozen plasma products must be stored with a core temperature of between -30 oC and -40oC throughout its shelf-life.

Platelets must be stored with a core temperature of between +20 oC and +24oC throughout its shelf-life.

4.4	Process Control

Temperature recording should be calibrated within +/- 0.5oC.

5	Functions

5.1	Operation

The system must be installed on both XXXX and YYYY sites by the manufacturer/supplier, who will provide documentation for Installation and Operational Qualification.

A calibration certificate (to UKAS Standards)  will be given for any equipment used to calibrate/install the refrigerator.

5.2	Control System requirements

There shall be at a core temperature probe, in a solution which simulates product (the formulation of this solution will be agreed with purchaser and ensure no growth of bacteria and will be contained in a secure manner).

Temperature shall be displayed digitally and graphically. Historic data should be readily available to easily see trends.

Alarm histories of each storage unit should be readily accessible with an audit trail associated with each event.

System may be able to be configured to allow grouping of fridge systems separately e.g. blood fridges, biochemistry fridges etc.

5.3	Power Failure/Recovery

In the event of a power failure, the system shall ensure no compromise of monitoring for a period of time to be specified by the manufacturer.

Data backup must ensure that in the event of hardware failure that the system integrity and temperature data integrity is not compromised.

5.4	Recovery methods

Manual restart to last state before power loss to the set points established by purchaser

Automatic restart to last state before power loss to the set points established by purchaser

5.5	Alarms/Warnings

Audible and visual alarms must be activated immediately (or without delay)on deviation of the product temperature from set limits.  If the temperature does not return to within set limits within 15 minutes the alarm will resound and every 15 minutes until temperature returns to within the set limits.

Audible and visual alarms may be activated on the door being opened for an extended period of time.

Audible and visual alarms should be activated on power failure

Alarm limits should be able to be modified by operator. Operator training is given to ensure the set points and alarms setting are not changed to out with the correct temperatures. Should the set point be changed accidentally then system will alarm as the set point and the alarm points are entered into it.

5.6	User interfaces

The control system shall include interfaces with the operator that ensures easy, safe and reliable operation.

The system should be compatible with Trust IT requirements.

The system must allow control of user access at different levels.

5.7	Environment

No special environmental conditions are required.

5.8	Physical conditions

Construction and function must ensure compliance with BS 4376

Must run on UK electrical supply

5.9	Environmental Classifications

No special environmental classification is required.

5.10	Biohazard level

The system shall be CE Marked

5.11	Cleaning Requirements

All surfaces must be accessible and be able to be wipe-cleaned

5.12	Sterilisation Requirements

No special sterilisation conditions are required

5.13	Intended Operating Environment

Operating Environment for the monitoring system will be within standard Laboratory room conditions, the probes will be in blood fridges, plasma freezers and platelet incubators located in the laboratory and around the hospital in various clinical areas.

- Requirements for Installation

6.1	Calibration

All temperature measuring and recording equipment must be calibrated with a tolerance of +/- 0.5oC and a calibration certificate provided for any equipment used to install/calibrate the system.

6.2	Deliverables

Deliverables identified in this document will be identified in advance and at the time of installation by the manufacturer.

  - 6.3      Documentation

Appropriate user manuals and schematics will be provided by the supplier

  - 6.4      Training

Appropriate training along with documentation will be provided by the supplier.

  - 6.5      Tools and Spares

Spares will be identified by the supplier.

  - 6.5      PPM/Calibration/Servicing/Re-Validation(Mapping)

Will be identified by the supplier and Trust as required